# IV Swinger 2

Step-by-step Construction:

Arduino Shield PCB Designs
   **FET module version**

Document Revision: 1.01  (24-Feb, 2022)			Chris Satterlee

Copyright (C) 2022  Chris Satterlee

IV Swinger and IV Swinger 2 are open source hardware and software projects.

Permission to use the hardware design is granted under the terms of the TAPR Open Hardware License Version 1.0 (May 25, 2007) - http://www.tapr.org/OHL

Permission to use the software is granted under the terms of the GNU General Public License v3 - http://www.gnu.org/licenses.

Current versions of the license files, documentation, PCB files, and software can be found at:

https://github.com/csatt/IV_Swinger

Step 1:	Understand the HW design / Choose Variant	5

Step 2:	Install software	5

Step 3:	Order PCB	5

Step 4:	Buy other parts	6

Step 5:	Gather / buy tools	7

Step 6:	Prepare for Soldering	8

Step 7:	1/4W resistors	9

Step 8:	IC sockets	10

Step 9:	Stacking connectors and female header	11

Step 10:	Screw terminal block	12

Step 11:	Filter capacitors	12

Step 12:	Bypass diode	13

Step 13:	Vertical shunt resistor	14

Step 14:	Vertical bleed resistor	15

Step 15:	FETs	16

Step 16:	Load capacitors	16

Step 17:	Optionally clean flux residue from PCB	17

Step 18:	Check for shorts	17

Step 19:	Insert ICs	18

Step 20:	Prepare load circuit wires	18

Step 21:	Make load circuit connections	20

Step 22:	Check off-PCB connections	21

Step 23:	Mate PCB with Arduino	21

Step 24:	Smoke test	22

Step 25:	Load Arduino sketch	22

Step 26:	Connect via IV Swinger 2 app	23

Step 27:	Apply resistor calibration	24

Step 28:	Sanity tests	24

Step 29:	Prepare for case and final assembly	25

Step 30:	Mark holes for Arduino standoffs	26

Step 31:	Mark holes for binding posts	27

Step 32:	Drill marked holes	27

Step 33:	Install binding posts	28

Step 34:	Install Arduino and PCB	29

Step 35:	Restore binding post connections	29

Step 36:	Drill USB connector hole	30

Step 37:	Make PV cables	31

Step 38:	Final test	31

Intro

This document contains the step-by-step instructions to build the FET module variant of IV Swinger 2.

These are the same instructions that are on www.instructables.com but do not include any of the photos or videos. The purpose of this document is primarily so you can print it out and check off completed steps and make other notes as you work on the project. You should use the Instructable to take advantage of the visual aids.

https://www.instructables.com/id/IV-Swinger-2-a-50-IV-Curve-Tracer

# Understand the HW design / Choose Variant

Please refer to the Instructable for this Step.

https://www.instructables.com/id/IV-Swinger-2-a-50-IV-Curve-Tracer

The remainder of this document assumes that you have chosen the following variant:

PV module version (FET)

https://github.com/csatt/IV_Swinger/raw/master/PCB/IV_Swinger_2_fet_mod/PDF/IV_Swinger_2_fet_mod_sch.pdf

https://github.com/csatt/IV_Swinger/raw/master/PCB/IV_Swinger_2_fet_mod/PDF/IV_Swinger_2_fet_mod_brd.pdf

If this is not correct, please find the step-by-step construction document for the version you did choose.  Many of the steps are similar, but the details differ so you need to use the correct document.

# Install software

Before spending time building the hardware, install the Arduino software and the IV Swinger 2 application on the laptop that you’ll be using.

- Install Arduino IDE:

https://www.arduino.cc/en/Main/Software

- Install IV Swinger 2 app:

https://github.com/csatt/IV_Swinger/releases

- Make sure both of the above come up before proceeding. If necessary, upgrade the OS on your computer

# Order PCB

Currently the PCB must be purchased from a manufacturing house that will actually fabricate it for your order. The downside of this is that you’ll probably have to buy more than you need. I have used the following two:

OSH Park:

https://oshpark.com

Made in USA

Cost: $25 for 3 PCBs (includes shipping)

Time: < 12 days to ship

PCBWay:

https://www.pcbway.com

Made in China

Cost: $5 for 10 PCBs + shipping ($16 DHL to CA)

Time: < 5 days to ship

Amazingly, I have put orders in to PCBWay on a Monday and had the boards in my hands in California on Friday.

I have shared this PCB design on PCBWay, and you can order it directly using the following link:

https://www.pcbway.com/project/shareproject/W112835ASP7_IV_Swinger_2_fet_mod_2021_10_25_1536_819e1ff4.html

https://www.pcbway.com/project/shareproject/W112835ASP7_IV_Swinger_2_fet_mod_2021_10_25_1536_819e1ff4.html

Alternately, you can order PCBs from OSH Park (or anywhere else) by uploading the ZIP archive of the Gerber files, which are found in the GitHub repository:

IV_Swinger/PCB/IV_Swinger_2_fet_mod/Gerber/*.zip

Soon, I hope to find someone who wants to sell individual PCBs on eBay (possibly in kits, that include all the other parts too).

# Buy other parts

The other necessary parts to build an IV Swinger 2 can all be purchased online from Amazon and Digi-Key.

FET PV module version BOM:

https://github.com/csatt/IV_Swinger/raw/master/PCB/BOM/fet_mod_BOM.pdf

The BOM has an Amazon link and a Digi-Key link at the bottom. The Amazon link is a “wish list” that can be used to populate your cart. Some of the items come in quantities larger (in some cases much larger) than needed to build a single IV Swinger 2. You may of course choose to find equivalents that are offered in smaller quantities. Also, many of the items are things that you may already have, so don’t necessarily just blindly order everything on the list.

The Digi-Key link is a pre-populated shopping cart. Again, you’ll want to check if you already have any of the items before ordering.

In both cases, it is possible (or probable) that certain items will go out of stock or be discontinued, so you’ll have to find suitable substitutions. Note that some of the Digi-Key items have ALT_n (e.g., ALT_1, ALT_2) in the “Customer Reference field. These should only be ordered if the primary version of the same part is marked as “backorder”.

Alternately, the following CSV file may be used to order from Digi-Key, Mouser or possibly other suppliers that support uploading a BOM spreadsheet. It includes the manufacturer, manufacturer’s part number, Digi-Key part number, and Mouser part number:

https://github.com/csatt/IV_Swinger/raw/master/PCB/BOM/DigiKeyMouser/fet_mod_DigiKey_Mouser.csv

https://github.com/csatt/IV_Swinger/raw/master/PCB/BOM/DigiKeyMouser/fet_mod_DigiKey_Mouser.csv

Also included below is the link to donate to the original Arduino developers. I donate $5 for each $10 Arduino clone that I buy. This is your choice, but I think it is the right thing to do.

Donate to Arduino.cc:
        https://www.arduino.cc/en/Main/Contribute

# Gather / buy tools

- Holding:

  - Vise

  - 3rd hand tool with magnifying glass

  - Tape (preferably Kapton, but Scotch ok)

  - Long/needle-nosed pliers

- Soldering:

  - Soldering iron (preferably temp controlled solder station)

  - Tip cleaner

  - Rosin core solder

  - Solder sucker or solder wick

- Cutting:

  - Wire cutter (flush cut)

  - Wire stripper

- Drilling:

  - Drill

  - 1/16" bit (pilot for 9/64")

  - 9/64" bit (standoffs)

  - 11/64" bit (pilot for 13/64")

  - 13/64" bit (binding posts)

  - 3/8" Forstner bit (preferred - USB cable hole)

    - Alternate: 1/8”, 3/16”, 7/32”, 1/4”, 9/32”, 5/16”, 11/32”, 3/8”, and 25/64” normal bits

- Other:

  - Digital Multimeter (DMM)

  - Small Phillips screwdriver

  - 9V battery

  - Sharpie

  - Ruler

  - Water spray bottle

# Prepare for Soldering

- Soldering NOTES:

- If you don't have a lot of soldering experience, read this:

https://learn.adafruit.com/adafruit-guide-excellent-soldering/common-problems

- Soldering components to the PCB is pretty mistake-proof, but doing it in the order described is recommended (shortest -> tallest).

- Some components have a correct and an incorrect orientation and some don’t matter. Pay attention to the instructions.

- I highly recommend using 63/37 0.031” (or 0.8mm) rosin core solder. Yes, it is 37% lead, but it is not a health risk for you (really), and environmentally insignificant when used by hobbyists. You’ll solder like a pro.

# 1/4W resistors

Resistors can be inserted in either orientation. It is very important to use the correct value for each, however.

- Solder 1/4W resistors to PCB:

- Insert all resistors before soldering. Tape down on front to hold in place OR bend leads slightly on back.

PV module version (FET) – 16 joints:

- R1  (150k):  _______

- R2  (7.5k):  _______

- R3    (1k):  _______

- R4    (1k):  _______

- R5   (22k):  _______

- R6  (390Ω):  _______  (390 ohms not k!)

- RF   (75k):  _______

- RG    (1k):  _______

- Flip board upside down and hold with vise or 3rd hand tool OR tape board to work surface. Solder all 16 leads _______

- Inspect with magnifying glass to make sure all joints are good and there are no solder bridges                          _______

NOTE: A solder bridge is ok between the ends of RF and RG

- Trim all leads _______

- Use multimeter to measure exact resistances of soldered resistors: 

With the PCB still upside down, measure the resistances with a DMM. The resistances (but unfortunately not the names) are marked on the back. Record the exact values of the ones marked with an asterisk (*) below – these values will be used later (Step 27:“Apply resistor calibration”). The others should just be close to their specified value (should be 1%, but don’t worry as long as it is < 10%) - the main point is to catch any mistakes you might have made.

PV module version (FET):

- R1  (150k):  _______  *

- R2  (7.5k):  _______  *

- R3    (1k):  _______

- R4    (1k):  _______

- R5   (22k):  _______

- R6  (390Ω):  _______  (390 ohms not k!)

- RF   (75k):  _______  *

- RG    (1k):  _______  *

# IC sockets

- Solder IC sockets to PCB – 24 joints:

- Insert three sockets before soldering. Tape down on front to hold in place.

- Make sure notch is on the LEFT end (TLV2462 and MCP3202) and the BOTTOM end (CPC1596) as marked on the PCB

- Flip board upside down and hold with vise or 3rd hand tool OR tape board to work surface and solder all 24 pins  
________

- Inspect with magnifying glass to make sure all joints are good

________

If you have opted not to use sockets, solder the ICs directly to the PCB instead of the sockets. Make sure dot is on the left end of the TLV2462 (pin 1). Make sure notch and dot are on the left end of the MCP3202 (pin 1). Make sure notch and dot are on the bottom end of the CPC1596 (pin 1).

# Stacking connectors and female header

- Solder stacking connectors and female header to PCB – 30 joints:

- Insert stacking connectors A1, A2, and A3 and female header FH. These connectors are symmetrical, so there’s no “backwards”. Tape down to hold in place.

      - A1 (10 pin):  ________

      - A2  (8 pin):  ________

      - A3  (8 pin):  ________

      - FH  (4 pin):  ________

NOTES: Stacking connector A4 is not needed

- Flip board upside down and hold with vise or 3rd hand tool OR tape board to work surface and solder all pins                                         ________

NOTE: the pins on A1, A2, and A3 that are actually used on the PCB are circled on the back of the PCB. Soldering the others provides physical support only.

- Inspect with magnifying glass to make sure joints are good and there are no solder bridges                          ________

# Screw terminal block

- Solder screw terminal block to PCB – 2 joints:

- Insert screw terminal block with the openings facing left. Tape down to hold in place.

      - J1:  ________

NOTE 1: The screw terminal block may be 2-pin or 3-pin. If a 2-pin block is used, insert it in the lower two holes and leave the top hole open.

NOTE 2: The screw terminal block may be omitted entirely, soldering the 18ga wire directly to the holes in the PCB (later).

- Flip board upside down and hold with vise or 3rd hand tool OR tape board to work surface and solder all joints                                         ________

- Inspect with magnifying glass to make sure joints are good and there are no solder bridges                          ________

# Filter capacitors

The small filter capacitors are not polarized, so it doesn’t matter which lead goes in which hole.

- Solder 0.1µF capacitors to PCB - 4 joints:

- Insert both capacitors before soldering. Bend leads on back to hold in place.

      - C3:  ________

      - C6:  ________

- Flip board upside down and hold with vise or 3rd hand tool and solder all four joints                                         ________

- Inspect with magnifying glass to make sure joints are good and there are no solder bridges                          ________

- Trim all 4 leads  _______

- Solder 2.2nF (2200pF) capacitors to PCB - 4 joints:

- Insert both capacitors before soldering. Bend leads on back to hold in place.

      - C4:  ________

      - C5:  ________

- Flip board upside down and hold with vise or 3rd hand tool and solder all four joints                                         ________

- Inspect with magnifying glass to make sure joints are good and there are no solder bridges                          ________

- Trim all 4 leads  _______

# Bypass diode

The purpose of the bypass diode (15SQ100, 15A/100V) is to protect the electronics in case the PV is connected to the IV Swinger 2 backwards.

- Solder bypass diode to PCB – 2 joints:

- Bend lead on striped end of diode around the diode so that it points in the same direction as the other lead.

- Insert leads as follows:

      - Pad D1 (cathode), striped end (top):  ________

      - Pad D2 (anode), non-striped end (bottom):  ________

- Flip board upside down and hold with vise or 3rd hand tool and solder both leads                                         ________

- Trim leads _______

- Re-flow/add solder on both leads  _______

(This is because leads are thick, and may not have heated well before trimming)

- Inspect with magnifying glass to make sure joints are good

________

# Vertical shunt resistor

The shunt resistor is oriented vertically on the PCB and should be soldered on at this point.

- Solder vertical shunt resistor to PCB – 2 joints:

- Bend one lead (either one) of the 5mΩ shunt resistor around the resistor so that it points in the same direction as the other lead:  ________

- Insert bent lead in lower hole and unbent lead in upper hole.  Tape in place.

      - SHUNT:  ________

- Flip board upside down and hold with vise or 3rd hand tool and solder both leads  ________

- Trim both leads _______

- Re-flow/add solder on both leads  _______

(This is because leads are thick, and may not have heated well before trimming)

- Inspect with magnifying glass to make sure joints are good

________

# Vertical bleed resistor

The bleed resistor is oriented vertically on the PCB and should be soldered on at this point.

- Solder vertical bleed resistor to PCB – 2 joints:

- Bend one lead (either one) of the 47Ω bleed resistor around the resistor so that it points in the same direction as the other lead:  ________

- Insert bent lead in right hole and unbent lead in left hole. Tape in place.

      - RB:  ________

- Flip board upside down and hold with vise or 3rd hand tool and solder both leads  ________

- Trim both leads _______

- Re-flow/add solder on both leads  _______

(This is because leads are thick, and may not have heated well before trimming)

- Inspect with magnifying glass to make sure joints are good

________

# FETs

- Solder FETs to PCB – 9 joints:

- Insert the three FETs (IRF540) with their metal side toward the right side of the PCB and the black side facing the left side of the PCB. Pin 1 is toward the top of the PCB. MAKE SURE THE METAL BACKS OF FET2 AND FET3 ARE NOT TOUCHING EACH OTHER! 
________

- Solder all 9 leads ________

- Trim all 9 leads ________

- Re-flow/add solder on all 9 leads ________

- Inspect with magnifying glass to make sure joints are good ________

# Load capacitors

- Solder load capacitors to PCB:

Module versions use 1000µF, 100V load capacitors.

These are polarized electrolytic capacitors, so orientation is important.

- Insert load capacitors in position.  Stripe side (shorter lead) goes to the right – this is the negative lead. Tape to hold in place.

    - C1    ________

    - C2    ________

- Flip board upside down and hold with vise or 3rd hand tool ________

- Solder all 4 leads  ________

- Trim all 4 leads _______

- Re-flow/add solder on all 4 leads  _______

(This is because leads are thick, and may not have heated well before trimming)

- Inspect with magnifying glass to make sure joints are good

________

# Optionally clean flux residue from PCB

Some people think it is important to clean off the flux residue from the PCB after soldering. It makes it looks nicer, but since the PCB sits on top of the Arduino, you don’t see the back anyway.

Functionally, it shouldn’t matter. The solder manufacturer Kester says this:

“Rosin flux residues are non-conductive and non-corrosive. Under normal circumstances they do not have to be removed from a printed circuit assembly. Rosin residue removal would be for cosmetic considerations. In an environment where the working temperature of the assembly will exceed 200°F the rosin residues will melt and become conductive, in these situations flux removal is required.”

If you do want to clean it off, see this Instructable: https://www.instructables.com/id/Cleaning-up-your-PCB/

# Check for shorts

Using the digital multimeter (DMM) set on the continuity check (beep), check that there is no continuity between the following:

Power to ground (mandatory):

- Left IC socket, pin 8 to pin 4
  OR

- Right IC socket, pin 8 to pin 4

Other (recommended):

- All “neighbor” pins or solder joints. None should indicate continuity, except the pairs circled in the pictures.

  - The idea is to find solder bridges that you didn’t see visually

Also:

- Measure the resistance between the metal backs of FET2 and FET3. It should be ~47Ω (resistance of RB)

# Insert ICs

Static electricity can destroy ICs. Take off your shoes and touch something metal connected to ground before handling them, if possible.

- Insert TLV2462 in left socket  _________

- Make sure dot is on the left end (pin 1)

- Legs may have to be bent inward slightly

- Insert MCP3202 in right socket  __________

- Make sure notch and dot are on the left end (pin 1)

- Legs may have to be bent inward slightly

- Insert CPC1596 in upper socket  __________

- Make sure notch and dot are on the bottom end (pin 1)

- Legs may have to be bent inward slightly

# Prepare load circuit wires

- Prepare load circuit wires:

- NOTE: This can be any stranded AWG 18 or AWG 16 insulated wire such as from a typical household extension/lamp cord or heavier speaker wire.  AWG 18 solid core is fine too (see alternate instructions below).

- "PV-": PV- (black) binding post to PV- screw terminal on PCB (J1)
“PV+”: PV+ (red) binding post to PV+ screw terminal on PCB (J1)

Stranded wire:

- Cut to length: 7 cm                                       PV- ________
PV+ ________

- Strip 1 cm on each end and twist strands                  PV- ________
PV+ ________

- Crimp cable ring connector on one end using pliers (or vise / ViseGrips / crimping tool)                                                  PV- ________
PV+ ________

- Heat crimp with the soldering iron and flow solder into strands
PV- ________
PV+ ________

- Heat the strands of the other twisted end and flow solder into the strands (i.e. "tin" it)                                  PV- ________
PV+ ________

Solid core wire (alternate):

- Cut to length: 8 cm                                       PV- ________
PV+ ________

- Strip 1 cm from one end and 2 cm from the other end                  PV- ________
PV+ ________

- Use long-nose pliers to bend the 2cm end into a loop                                                   PV- ________
PV+ ________

# Make load circuit connections

Refer to the drawing of off-PCB connections for this step.  These connections use the load circuit wires that were prepared in the previous step.

- Make binding post connections:

- Remove outer nuts and washers from threaded posts
________

- Insert threaded post of black side through the cable ring connector (or loop, if solid core wire is used) on load circuit wire:
  “PV-“

________

- Insert threaded post of red side through the cable ring connector (or loop, if solid core wire is used) on load circuit wire:
  “PV+“

________

- Put washers back on
________

- Put nuts on and tighten
________

- Make PCB connections:

- Loosen screw and insert the other end of the load circuit wire from the black binding post into the lower (-) hole of screw terminal J1 and tighten down the screw.
   “PV-“

_________

- Loosen screw and insert the other end of the load circuit wire from the red binding post into the other (+) hole of screw terminal J1 and tighten down the screw.
   “PV+“ _

________

# Check off-PCB connections

- Check off-PCB connections:

- Use the drawing of off-PCB connections and double-check that connections match the drawing.

__________

- Tug wires connected to screw-terminal blocks gently to make sure they are securely connected.

__________

# Mate PCB with Arduino

- Mate PCB with Arduino:

- Put tape on metal USB connector housing where PCB will touch it
________

- Line up stacking connector pins on bottom of the PCB with the corresponding connectors on the top of the Arduino and press the boards together, taking care not to bend any of the pins.  
________

# Smoke test

- Smoke test:

- Connect Arduino to laptop via USB

- Check for smoke    _______

- Check that Arduino yellow LED is blinking once per second (assuming that it’s still loaded with “Blink” sketch) _______

# Load Arduino sketch

- Load IV Swinger 2 Arduino sketch:

- Open Arduino application on your computer
________

- Find where the Arduino software looks for sketches:

  Arduino->Preferences->Sketchbook location

- Use your browser to go to:
https://raw.githubusercontent.com/csatt/IV_Swinger/master/Arduino/IV_Swinger2/IV_Swinger2.ino

https://raw.githubusercontent.com/csatt/IV_Swinger/master/Arduino/IV_Swinger2/IV_Swinger2.ino

- Right-click and use “Save As” to save IV_Swinger.ino to the Arduino sketchbook folder found above (make sure your browser doesn’t add an extension like .txt to the file name)

- Go back to the Arduino application and find the IV_swinger2.ino sketch using:

   File->Open

The Arduino application will inform you that IV_Swinger2.ino must be in a folder named IV_Swinger2 and it will offer to do that for you. Accept its kind offer.

- Click on arrow button or select “Upload” from “Sketch” menu
_________

- Check Arduino LEDs: Yellow LED should be blinking. This is not the same yellow LED that the Blink sketch controls.
_________

# Connect via IV Swinger 2 app

- Connect via IV Swinger 2 application:

- Open the IV Swinger 2 application  ________

- Verify that “Swing!” button text changes to RED and the message below it changes from “Not connected” to “Connected” (briefly, then disappears). The yellow LED should no longer be on.
_________

If not, pull down the “USB Port” menu and select the correct port.  If it isn’t obvious which one to select:

- Close the IV Swinger 2 application and disconnect the IV Swinger 2 USB cable from the laptop

- Re-open the IV Swinger 2 application (leave the cable disconnected)

- Pull down the USB Port menu and take note of the listed ports

- Connect the USB cable from the IV Swinger 2 hardware to the laptop

- Pull down the USB Port menu and select the port that is new to the list

# Apply resistor calibration

- Apply resistor calibration:

- In the IV Swinger 2 app, select “Resistors” from the “Calibrate” menu  ________

- Enter the values you measured and recorded in “Step 7: 1/4W resistors” above.

- Values are in ohms

________

# Sanity tests

- Sanity tests:

- “Nothing connected” test

- Click the “Swing!” button.  You should see an error dialog saying “ERROR: Voc is zero volts”
_________

- Battery test

Use 9V battery

- Strip both ends of two wires and screw one end of each into the side holes of the binding posts. If you happen to have a battery connector or holder with wires, use that. 
_________

- Connect the wire from the RED binding post to the positive terminal of the battery (you can either tape it or hold it with your thumb/finger)
_________

- Connect the wire from the BLACK binding post to the negative terminal of the same battery
_________

- Click the “Swing!” button.  You should get an IV curve that looks like the photo.
_________

- If you get an error dialog that says: “ERROR: Voc is zero volts” or “ERROR: Timed out polling for stable Isc”, check that you don’t have the battery backwards and that the wires are making good contact with the terminals.

# Prepare for case and final assembly

The acrylic baseball display case used for the IV Swinger 2 enclosure needs to have several holes drilled through it for attachments.

Case side definitions (see photo):

- Front: side with the USB connector

- Back: side opposite from front

- Left: side with binding posts

- Right: side opposite from left

- Bottom: side with Arduino

- Top: side above PCB

The case comes in two U-shaped halves:

- Base: Left / Bottom (with fins) / Right

- Lid: Front / Top / Back

All the attachments are made to the base half. The lid half has nothing attached to it, but does need a 3/8” hole in the front for the USB cable.

Care must be taken when drilling acrylic or else it will crack:

- Use a drill press if you have one

- Use vise (with rubber guards) to hold case

- Position so that the hole being drilled is close to the vise jaw

- Start with 1/16” pilot for all holes

- Drill slowly with light pressure

- Spray water on hole as it is being drilled to cool

- Use a Forstner bit to drill the 3/8” hole for the USB cable. Otherwise, you’ll have to start with 1/16” pilot and drill incrementally larger holes until you get to 3/8” (actually 25/64”)

# Mark holes for Arduino standoffs

IMPORTANT: For this step and the next three, look straight down with one eye when making the Sharpie dots (the plastic distorts/refracts if you look at an angle, and you’ll miss the mark).

- Mark holes for Arduino standoffs:

- Attach four 15mm standoffs to Arduino:

- Unplug the USB cable from the Arduino
_______

- Carefully remove the PCB from the Arduino
_______

- Insert threaded/male end of each standoff through its hole in the Arduino from the back
________

- Screw nuts onto the threaded ends of the standoffs on the front of the Arduino – hold the nut with your finger and turn the standoff to tighten it. Use pliers to tighten more.

NOTE: The hole nearest the Arduino reset button doesn’t have room for a nut
________

- Place the Arduino in position, standing on its standoffs (including the one without a nut). The Arduino should be touching the right side of the case, with the USB connector facing the front. The single fin should be facing toward you so the fins look like a “Y”. See photo.
________

- PUT LID ON THE CASE. This is important because the fit is very tight!
________

- Turn the case over and look at it from the bottom. The Arduino will probably stay in place, but you can make sure by squeezing the front and back together with the hand you’re holding it with. Use a Sharpie to mark the centers of the four holes. 
________

- Remove the lid from the case and remove the Arduino
________

# Mark holes for binding posts

- Mark holes for binding posts:

- Remove top nuts, washers, cable rings, and bottom nuts from the binding posts. Remove the black plastic backing plate.
________

- Hold the plastic backing plate in position on the inside of the left side of the case. It should be about 1mm from the front inner edge of the case and about 1mm from the bottom.
________

- Use Sharpie to mark the centers of the two holes
________

# Drill marked holes

- Drill 6 marked holes:

- Use something pointy to make an indentation in the middle of each of the Sharpie marks. The tip of the Forstner bit is perfect for this, but you can also use a needle or the tip of an X-acto blade (poke and twirl). This will keep the drill bit centered when you start drilling the hole.
________

- Drill 1/16” pilot holes
________

- Switch to 9/64” bit and re-drill all holes
________

- Enlarge holes for binding posts:

- Switch to 11/64” bit and re-drill the binding post holes 
________

- Switch to 13/64” bit and re-drill the binding post holes one more time 
________

- Clean up case:

- Remove burrs around holes with X-acto knife or your fingernails
________

- Wash case off and dry
________

# Install binding posts

- Install binding posts:

- Insert the binding posts through their holes with the RED terminal toward the TOP of the case
________

- Slide backing plate over the posts on the inside of the case
________

- Thread nuts on the posts and tighten down
________

# Install Arduino and PCB

- Install Arduino (without PCB) in case:

- Attach the one Arduino standoff that won’t have a nut onto the bottom of the case with an M3 screw
________

- Insert the Arduino, put the lid on the case, and screw down the other three standoffs with M3 screws. TIP: start all screws before tightening any of them.
________

- Remove the lid
________

- Mate PCB back onto Arduino:

- Load circuit wires should still be screwed to PCB.  If not, insert them back to into their correct screw terminal block openings and tighten them down.
________

- Line up stacking connector pins on bottom of the PCB with the corresponding connectors on the top of the Arduino and press the boards together, taking care not to bend any of the pins. 
________

# Restore binding post connections

- Restore connections to binding posts:

- Restore connections as before, following the off-PCB connections drawing. Tighten nuts securely.
_________

# Drill USB connector hole

- Drill USB connector hole:

- Put the lid on the case
_________

- Make indentation in the exact center of the USB connector using the tip of the Forstner bit (or whatever pointy thing you used for the other drill-starting indentations). NOTE: it is very important that this hole is precisely centered. You need to look at it from all four directions before making the indentation since the refraction through the plastic distorts the apparent position (you’ll see what I mean as soon as you turn it 90 degrees).
_________

- Use 3/8” Forstner bit to drill the hole

- Drill slowly, spraying with water often

- Reduce pressure when hole is getting close to “punching through”

- Alternative to Forstner bit is to use following succession of normal bits:

  - 1/16”, 1/8”, 3/16”, 7/32”, 1/4”, 9/32”, 5/16”, 11/32”, 3/8”, 25/64”

_________

- Clean up the edge of the hole with X-acto knife or your fingernail
__________

- Wash lid off and dry ________

- Put lid on and insert the USB cable to make sure it fits ________

- If it doesn’t, try loosening the Arduino standoff screws. This might give you enough “play” to get the connector in. Then, with the connector still in, re-tighten the screws

- If that isn’t enough, you may have to enlarge the hole with a round file or some other way

# Make PV cables

- Make PV cables:

To connect to a standard PV module, you need cables with MC4 connectors.

It is not necessary to use the same heavy gauge cable that is used in a rooftop solar installation (and on the modules themselves), assuming you only need them to be a few feet long. The nice thing about the binding posts is that you can easily swap cables with longer or shorter ones depending on the situation. The main reason for longer cables would be so the laptop and IV Swinger 2 can be in a shady spot away from the panel. These instructions intentionally do not specify the length or type of the PV cables because it is so dependent on the usage.

If you decide that shorter cables are OK, you can just use the same load circuit wire that you used for the internal load connections.  The only tricky part is that crimping the MC4 connectors onto smaller wire gauge doesn't really work - you need to solder them on. You also should use solder to tin the bare ends that insert into the binding posts so they are more durable.

The downside to the binding posts is that it is possible to connect the wrong cable to the wrong post. The bypass diode(s) protect against this, but it's still a good idea to make it as foolproof as possible. Put some red tape around the one that connects to the red binding post and some black tape around the one that connects to the black binding post.

The cable with the female MC4 connector connects to the RED binding post.

The cable with the male MC4 connector connects to the BLACK binding post.

# Final test

Your IV Swinger 2 is now complete!

Repeat the tests you did in “Step 28: Sanity tests” to make sure everything got hooked back up correctly.

You may now test it with a real PV module.

If accuracy is important to you, see the IV Swinger 2 User Guide for instructions on how to perform a calibration. There is also a Help dialog available from the Calibrate menu in the application.